OPERATION MANUAL

INDEX

1. Introduction     ………………………………… (1)

2. Specifications …………………………………… (1)

3. Front panel & operations …………………………(2)

4. Operation …………..………………………..…….(3)

5. Accessories & Maintenance ……………..……….(4)

INTRODUCTION

This series power supply is a newly developed model with high stability and accuracy. Its unique block-like structure makes it easily to be extended from single channel to multi-channel. It is an ideal instrument for R&D departments, colleges and factories.

SPECIFICATIONS

| Line Voltage | 105-130VAC, 50Hz-60 Hz 200-240VAC, 50Hz-60Hz |
| --- | --- |
| Input Current (full load) | 110V | 0.8A | 1.2A | 1.5A |
|  | 220V | 0.4A | 0.6A | 0.9A |
| Output Voltage Range | 0-30V | 0-30V | 0-30V |
| Output Current Range | 0-2A | 0-3A | 0-5A |
| Voltage Display Accuracy | ±0.5% |
| Current Display Accuracy | ±0.5% |
| Efficiency | ＞85% |
| Voltage Regulation |
| Load Regulation(10-100%) | 50mV | 50mV | 50mV |
| Line Regulation(200-240VAC) | 20mV |
| Ripple & Noise(P-P) | 30mV | 50mV | 50mV |
| Current Regulation |
| Load Regulation(10-100%) | 20mA |
| Line Regulation(200-240VAC) | 20mA |
| Ripple & Noise(P-P) | 20mA |
| Others |
| Operating Ambient | -10℃ to + 60℃/ 30%RH to 90% RH |
| Size | 7(W) x 16(H) x 22(D)cm |
| Weight | 1.35kg |

1

FRONT PANEL & OPERATION

Functions of the controls on front panel

(1)  3 digit LED display of voltage.

(2)  3 digit LED display of current.

(3)  CV (Constant Voltage) indicators. If the current passing through the load is lower than the preset value of current, the power supply is working in Constant Voltage state, and CV indicator lights.

(4)  Power Switch.

(5)  CC (Constant Current) indicators. If calculated with Ohm’s law, the current passing through the load would be larger than preset value of current, the power supply will work in Constant Current state, CC indicator lights, and the actual current is limited to the preset value of current.

(6)  Output terminals.

(7)  Current adjusting knob.

(8)  Voltage adjusting knob.

2

OPERATION

Connection of output ports:

Plugging test leads into port “+” and “-” only, it will work in floating mode. And you can also get plus or minus voltage by connecting “+” or “-” with “GND”.

Setup Constant Voltage:

Keep test leads open, adjust voltage knob to get desired voltage.

Setup Constant Current:

Turn voltage and current knobs CCW to their small value positions, short circuit the test leads, adjust current knob to get a desired current.

ACCESSORIES:

Power Cord                                 1pc

User’s Manual                               1pc

Spare Fuse (1.5A, 250 VAC, Ø5 x 20mm)               1pc

3

CAUTION:

1、If the Power Supply can’t be switched on, and line power is confirmed with no problem, the fuse of Power Supply may be broken. Switch off the Power Supply, disconnect the power cord, and change the fuse.

2、When working in Constant Voltage state，if output voltage lower than what was preset and CC indicator lights, the instrument automatically turn to Constant Current state. You must check the load or increase the output current.

3、When working in Constant Current state, if output current lower than what was preset and CV indicator lights, the instrument automatically turn to Constant Voltage state. You must check the load or increase the output voltage.

4、When the instrument is unstable in Constant Voltage state，Probably the line voltage is under 90% of the rated value. If the problem is not caused by the line voltage, contact your nearest dealer.

4

# Text boxes

SWITCHING DC POWER SUPPLY